# Optimisation Workflow Manual

## Overview

This guide provides a comprehensive, step-by-step workflow for using the iolite Optimiser. It combines the Single Pulse Response (SPR) analysis with the Method Optimiser to ensure your laser ablation parameters are perfectly tuned for your specific hardware and analytical goals.

## Phase 1: Determine Washout (SPR Tab)

Before optimizing spot size or speed, you must know your system’s washout time to prevent signal mixing.

### Step 1.1: Load SPR Data

- Ensure your iolite session contains a file with single laser pulses (e.g., a “Test” or “Tune” line created with a low repetition rate, typically 1 Hz).

- Open the Advanced Spot Optimiser plugin.

- Go to the SPR Tab.

- Click Reload Data from iolite to ingest the current selection.

### Step 1.2: Select & Optimize

- Select Channel: Choose a high-abundance isotope (e.g., U238, Th232) from the dropdown.

- Verify Detection: Look at the SPR Plot. Valid pulses should be marked with an orange X.

  - Too much noise? Increase Peak Cutoff (uncheck Auto) or Min. Distance.

  - Missed peaks? Decrease Peak Cutoff.

- Clean Data: If you see double-peaks or artifacts, click on them in the plot to Exclude them from the calculation.

### Step 1.3: Apply Washout

- Review the FW 0.1M (10%) statistics in the left panel.

- Decide which metric to use:

  - Average: Best for stable, well-behaved systems.

  - Max: Conservative; ensures zero signal overlap even for the slowest washout event.

- Click Apply Average (or Max) to Optimiser. This sends the value directly to the Optimiser tab.

## Phase 2: Optimiser Setup (Optimiser Tab)

### Step 2.1: Hardware Configuration

- Switch to the Optimiser Tab.

- Click the Settings button.

- ICP-MS: Select your Manufacturer and Model.

  - Custom? Choose “Custom Model” and enter your Min Dwell Time (e.g., for TOF or generic Quad).

- Laser: Select your Platform and Cell Type.

  - Custom? Choose “Custom Laser” and manually limit the Max Rep Rate and Max Speed.

### Step 2.2: Load Sensitivity Data

- Ensure your iolite session has a representative ablation signal (e.g., a “Tuning” line or standard block).

- Click Reload Data from iolite.

- Dwell Time Check:

  - If your data lacks dwell time metadata, a dialog will appear.

  - Choose Set to Global (e.g., 10ms) or enter Individual times per channel to match your instrument method.

- Region Selection:

  - Check the Signal Plot.

  - Ensure the Background (Blue) region covers the gas blank.

  - Ensure the Signal (Red) region covers the steady ablation signal.

  - Use Auto Select Regions if they are incorrect.

## Phase 3: Run Optimization

### Step 3.1: Define Analytical Goals

- Mode: Select your analysis type (usually Imaging or Line).

- Washout: Verify this matches the value you applied from the SPR tab (Phase 1).

- Parameters:

  - Pulses per Dwell: Set the stability requirement (e.g., 10 shots/dwell).

  - Target SNR: Set the desired quality (e.g., 10 Sigma).

  - Avoid Gaps: Check this for Imaging to ensure full surface coverage.

### Step 3.2: Interpret Results

The software calculates the optimal settings automatically.

- Review Settings: Look at the Optimised Settings panel for your ideal Spot Size, Rep-Rate, and Speed.

- Check Constraints: Look at the Results Table

  - Orange Rows: These isotopes cannot meet the Target SNR even at the maximum allowable settings.

  - Blue Rows: These are limited by your hardware speed (Min Dwell).

- Refine:

  - Spot Size too big? Reduce the Target SNR or Pulses per Dwell.

  - Scan too slow? Uncheck Avoid Gaps (if acceptable) or increase Spot Size manually.

### Step 3.3: Manual Overrides & Advanced Logic

- Force Spot Size: If you strictly need a specific beam diameter (e.g., 25µm), enter it in the Spot Size spinbox. Valid Rep-Rate and Speed will be re-calculated for that fixed spot.

- Min SNR (Hardware Differences):

  - Quadrupoles: Channels failing SNR are set to the minimum dwell to save time.

  - MC/TOF: Channels failing SNR are flagged (Orange) but dwell times remain fixed as they are simultaneous detectors.